## Slide 1: Analysis of the Radiation Monitoring System's Initial Data from the GEM Detectors at the CMS Experiment
M. Misheva, P. Iyadjiev1, I. Vankov1, L. Dimitrov1,
G. Mitev1, L. Ratchev1
on behalf of the CMS Muon Group
11th International Conference of the Balkan Physical Union
28 August – 1 September 2022, Belgrade, Serbia
BPU11 CONGRESS
1 Institute for Nuclear Research and Nuclear Energy, BAS

## Slide 2: Outline
Challenges
GEM Detectors
Radiation Monitoring System
System structure
Initial Data before RUN 3

## Slide 3: CMS Experiment somewhere at the beginning of Long shutdown 2 (2019)
Endcap
Barrel
Hadron Calorimeter
Electromagnetic calorimeter
Tracker
Magnet
Muon chambers
Endcap nose
CMS Muon System – 3 detection technologies
Drift tubes (DTs)
Cathode Strip Chambers (CSCs)
Resistive Plate Chambers (RPCs)

## Slide 4
Challenges
LHC (RUN 3) – instantaneous luminosity close to 2x1034 cm-2s-1
HL-LHC (RUN 4 ….) – instantaneous luminosity close to (5-7.5)x1034 cm-2s-1
HL-LHC background – 5 x rates and 6 x total doses with respect to LHC
  Exceed the primary design tolerances of different components of the detectors - new assessment of the detectors electronics longevity, operation and performance
We are here now
CMS Muon System goal – maintain excellent triggering, ID and measurement of muons during HL-LHC

## Slide 5
Barrel
18 short & 18 long super-chambers (SCs) per endcap
short
long
long
Muon endcap Station 1
Endcap
Interaction Point
CMS endcap muon detectors:
CSC + RPC covering 0.9 < η < 1.6
Only CSC covering 1.6 < η < 2.4
GE1/1 + CSC covering 1.55 < η < 2.18
CMS Forward Muon System Upgrade – Phase I
}
Station 1
}
Station 2
}
Station 3
}
Station 4
10o
GE1/1 are new forward muon detectors to handle together with CSC at Station 1 with the most difficult region - high background trigger and readout rates, and limited bending
1 SC is made of 2 back to back trapezoidal-shaped triple-GEM detectors

## Slide 6
GE1/1 detectors
GE1/1 Installation
July-Oct 2019  Negative endcap & July-Sept 2020  Positive endcap
High rate capability, up to O(MHz/cm2)
Efficiency > 97%
Space (time) resolution ~ 300µm (8ns)
Gas mixture Ar/CO2 70/30
CMS Technical Design Report for the Muon Endcap GEM Upgrade
Microscope picture of a GEM foil
Schematic view of a GEM chamber
triple-GEM technology
assembly of the FE electronics
cooling circuit installation
metal enclosure cover attachment
SC assembly

## Slide 7
Radiation Monitoring
System in
GEM detectors

## Slide 8: Radiation Monitoring System in GEM Detectors
p-i-n diodes RadFETs
t°
REM 130
REM 250
BPW 34S
LBSD Si-1
NTC 10k
1 k
Each RadMon unit contains:
RadFETs (Radiation sensitive Field Effective Transistors) – absorbed dose (REM250, REM130)
 p-i-n diodes – 1 MeV neutron equivalent fluence (BPW34S, LBSD Si-1)
 Thermistor
 1kΩ resistor
Function | Type | Device | Operating range | Sensitivity/ Resolution | I read
Total Dose Sensor (high doses) | RadFET | REM 250 | A few 10-1 Gy up to > 2x104 Gy | 20 mV/Gy (initial) | 160 µA
Total Dose Sensor (very high doses) | RadFET | REM 130 | A few Gy up to > 2x105 Gy | 3 mV/Gy (initial) | 160 µA
1 MeV n eq. Fluence Sensor (high sensitivity) | p-i-n diode | LBSD Si-1 | 1010 cm-2 up to 2x1012 cm-2 (almost linear) | 2.1x108 cm-2 /mV | 10 mA
1 MeV n eq. Fluence Sensor (low sensitivity) | p-i-n diode | BPW34S | 2x1012 cm-2 up to 4x1014 cm-2 (linear) | 1x1010 cm-2 /mV | 1 mA
Temperature sensor | Thermistor | NTC 10k | -55 oC up to 125 oC | 0.1 oC | 10 µA
Line checking | Resistor | 1 k |  | 1 % | 160 µA
RadFETs and p-i-n diodes have several
advantages:
  Small detector sizes
  Immediate (online) readout
  Minimal power requirement
A monitoring system, designed to measure the radiation dose and particle fluence around the GEM detectors, has been developed and installed in GE1/1 chambers.
The Radiation Monitoring (RadMon) System consists of 12 RadMon units with active radiation sensors.

## Slide 9
Every RadMon unit has been inserted in metal bracket fixed on the GE1/1 cooling module (long chambers, wide region)
6 RadMon units with active radiation sensors are installed in 6 GE1/1 long detectors in each endcap (totally 12 sensors)
Splitter 1
Splitter 2

## Slide 10: Dose and particle fluence measurement
The voltage drop across all radiation
sensors is proportional to the measured radiation/particle fluence magnitude.
ΔU=Ui – U0 [V] (Si-1 or REM250)
before
Time
LHC start
after
For the RadFETs the relation between the gate threshold voltage shift ΔVth and the radiation dose D can be best approximated by ΔVth = A×DB, (resp. D = (ΔVth/A)1/B) . The coefficients A and B depend on the RadFET type as well as on the measured dose range.
The shift of the p-i-n diodes forward voltage ΔVF is proportional to the 1 MeV neutron equivalent fluence Φ [cm-2]. The relation is generally linear – Φ = cΔVF, where c depends of the diode type.
Irradiation with high energetic hadrons can provoke accumulation of local defects in the RadFETs/p-i-n diodes.

## Slide 11: Temperature Dependence
The uncertainty of threshold voltage of sensors is ±0.0018V considering both the precision of the ADC in the controllers and data averaging
The uncertainty of thermistor measurements is estimated to be ±0.22°C
No presence of magnetic field
The plots contain the averaged data per hour for the LBSD Si-1/RadFET (Rem250) and thermistor data which are integrated in RadMon 3 unit installed in the bottom sector of the CMS Negative endcap.

## Slide 12
Dependence on the CMS
Magnetic Field
The plots contain the averaged data per hour for the LBSD Si-1/RadFET (Rem250) and the CMS Magnetic field during one week from the “warming up period of the CMS Magnet” performed in October 2021.
The uncertainty of threshold voltage of sensors is ±0.0018V considering both the precision of the ADC in the controllers and data averaging
The temperature during the presented period fluctuated within 0.3°C
The p-i-n diode and RadFET are integrated in RadMon 6 unit installed in the top sector of the CMS Negative endcap.

## Slide 13
Daily average threshold voltage values for the p-i-n diode (Si-1) and RadFET (Rem250) active sensors in relation to the temperature and the CMS magnetic field during the last ten months of LS2 (Aug 1 2021 until May 26 2022 inclusively) and the beginning of the LHC Commissioning period (27th May 2022 until 20th June 2022)
The uncertainty of threshold voltage of radiation sensors is evaluated to be ±0.0004V due to data averaging and precision of the ADC integrated in the controllers while the uncertainty of thermistor measurements is estimated to be ±0.1°C.
The increase of forward Uth(Si-1) at given forward current in magnetic field observed in p-i-n diodes can be attributed to magneto-resistance as well as to the increased charged carrier recombination [1]. The observed change of Uth(REM250) in magnetic field is due to similar effects as in p-i-n diodes.
[1] I. Mandic´, F. Ravotti, Member, IEEE, M. Glaser, I. Serˇsa, J. Hartert, S. Franz, V. Cindro, I. Dolenc, A. Gorisek, G. Kramberger, and M. Mikuz,
“The Effect of Magnetic Field on Readout of Diodes Used as NIEL Counters “, IEEE Transactions on Nuclear Science ( Volume: 58, Issue: 3, June 2011)

## Slide 14: Conclusion
Acknowledgements
This work was supported by research grants from National Roadmap for Research Infrastructure 2020-2027, CERN, Ministry of Education and Science (MES), Bulgaria (D01-374/18.12.2020) and National Science Fund of Bulgaria (BNSF), MES (KP-06-CERN/2 /14.12.2018)
The threshold voltage values of the active radiation sensors are well determined before the start of LHC RUN3:
  Calibration (offset) of the voltage threshold for each sensor has been determined at the CMS magnetic field
  Temperature correction is required to achieve lowest possible uncertainty
The radiation monitoring system of GEM is fully tested to measure the total dose and neutron flux in the area of the GE1/1 chambers installed at Station1 in the region of the endcaps.

## Slide 15: Thank You